## Slide 1: Regulatory Approval of New Devices
SCMR 2015 Interventional CMR Workshop
Technology for Interventional cMRI
Steve Wedan

## Slide 2: Regulatory Approval
Vital to help ensure true safety and effectiveness of new devices
Tool to help:
  Guide us to collective body of safety knowledge: standards
  Hold us accountable for careful and methodical design, manufacture, and testing of devices
  Force us to make a proof argument for the safety and effectiveness of new devices

## Slide 3: Regulatory Approval
Vital to help ensure true safety and effectiveness of new devices
Tool to help:
  Guide us to collective body of safety knowledge: standards
  Hold us accountable for careful and methodical design, manufacture, and testing of devices
  Force us to make a proof argument for the safety and effectiveness of new devices

## Slide 4: Where are the iCMR devices?
We have seen presentations and posters for years that show the feasibility of iCMR devices
Yet very few have made it to clinical practice
Regulatory approval process is extremely demanding
  MR safety testing can be tricky (applicable standards not always available)
  Standards based patient safety requirements can make MR compatibility solutions challenging

## Slide 5: Example: RF Heating of a Catheter
Non-magnetic catheter with no RF heating mitigation technology
ASTM F2182 Phantom
Saline 4.7 mS/cm
SAR = 4 W/kg

## Slide 6: Example: RF Heating of a Catheter
ASTM F2182 Phantom
Saline 4.7 mS/cm
SAR = 4 W/kg
Anatomical Trajectory
Non-magnetic catheter with no RF heating mitigation technology
ΔT = 2.3 °C

## Slide 7: Example: RF Heating of a Catheter
ASTM F2182 Phantom
Saline 4.7 mS/cm
SAR = 4 W/kg
Anatomical Trajectory
Catheter Cable Attached
Non-magnetic catheter with no RF heating mitigation technology
ΔT = 2.3 °C

## Slide 8: Example: RF Heating of a Catheter
ASTM F2182 Phantom
Saline 4.7 mS/cm
SAR = 4 W/kg
Anatomical Trajectory
Catheter Cable Attached
Non-magnetic catheter with no RF heating mitigation technology
ΔT = 2.3 °C
ΔT = 3.0 °C

## Slide 9: Example: RF Heating of a Catheter
ASTM F2182 Phantom
Saline 4.7 mS/cm
SAR = 4 W/kg
Anatomical Trajectory
Catheter Cable Attached
Improved Test Trajectory
Non-magnetic catheter with no RF heating mitigation technology
ΔT = 2.3 °C
ΔT = 3.0 °C

## Slide 10: Example: RF Heating of a Catheter
ASTM F2182 Phantom
Saline 4.7 mS/cm
SAR = 4 W/kg
Anatomical Trajectory
Catheter Cable Attached
Improved Test Trajectory
Non-magnetic catheter with no RF heating mitigation technology
ΔT = 2.3 °C
ΔT = 3.0 °C
E

## Slide 11: Example: RF Heating of a Catheter
E fields in the body
are complex

## Slide 12: Example: RF Heating of a Catheter
ASTM F2182 Phantom
Saline 4.7 mS/cm
SAR = 4 W/kg
Anatomical Trajectory
Catheter Cable Attached
Improved Test Trajectory
Varying Insertion Depths
Non-magnetic catheter with no RF heating mitigation technology
ΔT = 2.3 °C
ΔT = 3.0 °C

## Slide 13: Example: RF Heating of a Catheter
ASTM F2182 Phantom
Saline 4.7 mS/cm
SAR = 4 W/kg
Anatomical Trajectory
Catheter Cable Attached
Improved Test Trajectory
Varying Insertion Depths
Non-magnetic catheter with no RF heating mitigation technology
ΔT = 2.3 °C
ΔT = 3.0 °C
ΔT = 14 °C

## Slide 14: Example: RF Heating of a Catheter
Non-magnetic catheter with no RF heating mitigation technology
ASTM F2182 Phantom
Saline 4.7 mS/cm
SAR = 4 W/kg
Catheter Cable Attached
Improved Test Trajectory
Varying Insertion Depths
ΔT = 2.3 °C
ΔT = 3.0 °C
ΔT = 14 °C

## Slide 15: Example: RF Heating of a Catheter
Non-magnetic catheter with no RF heating mitigation technology
ASTM F2182 Phantom
Saline 4.7 mS/cm
SAR = 4 W/kg
Catheter Cable Attached
Improved Test Trajectory
Varying Insertion Depth
ΔT = 2.3 °C
ΔT = 3.0 °C
ΔT = 14 °C
ΔT = 22 °C

## Slide 16: Example: RF Heating of a Catheter
ASTM F2182 Phantom
Saline 4.7 mS/cm
SAR = 4 W/kg
Catheter Cable Attached
Improved Test Trajectory
Varying Insertion Depth
Tissue Mimicking Gel
Non-magnetic catheter with no RF heating mitigation technology
ΔT = 2.3 °C
ΔT = 3.0 °C
ΔT = 14 °C
ΔT = 22 °C

## Slide 17: Example: RF Heating of a Catheter
ASTM F2182 Phantom
Saline 4.7 mS/cm
SAR = 4 W/kg
Catheter Cable Attached
Improved Test Trajectory
Varying Insertion Depth
Tissue Mimicking Gel
Non-magnetic catheter with no RF heating mitigation technology
ΔT = 2.3 °C
ΔT = 3.0 °C
ΔT = 14 °C
ΔT = 22 °C
ΔT = 72 °C

## Slide 18: Example: RF Heating of a Catheter
A better way to test a new device for MR safety involves methods from ISO TS 10974:
“Assessment of the safety of magnetic resonance imaging for patients with an active implantable medical device”
This may soon be required, as the TS becomes a standard

## Slide 19: Example: Physiological Signal Filter
EP Amplifier
fc = 680 kHz
-85 dB @ 64 MHz
L = 75μH
C = 1600 pF
C = 1600 pF

## Slide 20: Example: Physiological Signal Filter
EP Amplifier
fc = 680 kHz
-85 dB @ 64 MHz
L = 75μH
C = 1600 pF
C = 1600 pF

## Slide 21: Example: Physiological Signal Filter
EP Amplifier
fc = 680 kHz
-85 dB @ 64 MHz
L = 75μH
C = 1600 pF
C = 1600 pF
Ileak

## Slide 22: Example: Physiological Signal Filter
EP Amplifier
fc = 680 kHz
-85 dB @ 64 MHz
L = 75μH
C = 1600 pF
C = 1600 pF
Ctotal < ~100 pF

## Slide 23: Where are the iCMR devices?
They are coming, but slowly
Part of the problem:
People tend to underestimate the challenge

## Slide 24: We need to help each other.1
Researchers shouldn’t necessarily be experts in regulation and standards
But experts in regulation and standards do exist
We are happy to help
  We are not selling a test service, but we can offer a bit of guidance
  We can help you:
    Understand the safety issues
    Do the right tests the right way
    Properly scope out what yet needs to be done
1 Lederman R., Schaeffter T., Email to SCMR iCMR Workshop speakers, 2015

## Slide 25: Regulatory Approval of New Devices
SCMR 2015 Interventional CMR Workshop
Technology for Interventional cMRI
Steve Wedan (CEO): steve.wedan@imricor.com
Tom Lloyd (VP Research): tom.lloyd@imricor.com