## Slide 1: PRiME (Physiological Recording in MRI Environment)
John Kakareka
Signal Processing and Instrumentation Section,
Division of Computational Bioscience, Center for Information Technology
NIH
Conflict of Interest Statement
  I have no financial relationships regarding the PRiME system
  NIH does not have any patents or licenses on the PRiME system
  NIH does not officially recommend or endorse any company or product mentioned

## Slide 2: PRiME Overview
Supports physiological recording of 6-lead ECG and 2 invasive blood pressure (IBP) channels
This is not a marketed system, but rather design documents on how sites can assemble a non-significant risk medical device.
Please adhere to local clinical and research standards if you plan to use these devices on patients.
Benchmarks
      >24 pediatric patients at Children’s National Medical Center in Washington, D.C.
      >6 adult patients at NIH
      numerous animal studies at NIH

## Slide 3: PRiME 1st Generation System Components
  Physiological Signal Acquisition – resides on patient table
    Designed specifically to reduce noise (i.e., gradient, RF pulse sequence) generated in an MRI environment
    Uses fiber optic cables to transfer digitized signals to control room
  Control System and Adaptive Filtering – resides in control room
    Uses gradient control signals from MRI to adaptively cancel gradient induced noise
      Real-time FPGA-based system designed using National Instruments Labview programming software
    Allows for display of signals; user control of system settings
    Provides trigger based on ECG or IBP signals which can output to MRI system
      Can be used for gated MRI sequences
  Physiological Signal Converter – resides in either control room or X-ray room
    Outputs transducer-level analog signals for integration with commercial recording systems (e.g., Siemens SENSIS, GE Maclab)

## Slide 4: PRiME Example of Gradient Noise Removal
Figure 1. Example of removal of MRI gradient-induced noise in a volunteer during a real-time imaging sequence. The top plot is the right arm electrode after acquisition and the bottom plot is same signal after the adaptive filter

## Slide 5: PRiME Technology Transfer for Early Adopters
  All schematics, printed circuit board layouts, mechanical designs, bill of materials and software code will be made available at no cost with no licenses required
  Guidelines on how to assemble and test a system will be provided
  Information enabling an electronics assembly house to assemble all three components
  http://nhlbi-mr.github.io/PRiME/
  Total estimated cost (including all hardware and software): $20,000
Figure 1. Physiological Acquisition System (black box) on IV Pole
Figure 2. Control System and Adaptive Filter System located in the MRI control room
Figure 3. Physiological Converter System located in MRI control room

## Slide 6: PRiME Acknowledgements
NHLBI, Cardiovascular Intervention Program
  Dr. Tony Faranesh
  Dr. Robert Lederman
Children’s National Medical Center
  Dr. Kanishka Ratnayaka
CIT, Signal Processing and Instrumentation Section
  Tom Pohida
  Randy Pursley
  Jessica Crouch