## Slide 1
New Designs & Sequences for
Device Tracking
Ehud J Schmidt1
With Contributions from: Wei Wang1, Charles L Dumoulin2 , Barret Daniels2, Zion TH Tse3, Israel Byrd4, Jeffrey Schweitzer4, Ronald Watkins5, Raymond Y Kwong6, William G Stevenson6 , Ravi Seethamraju7, Henry Halperin8, Kamal Vij9, Charles Cunningham10 Graham Wright10
1Radiology & 6Cardiology, Brigham and Women’s Hospital, Boston, MA
2Radiology, Cincinnati Children’s Hospital Medical Center
3Engineering, The University of Georgia
4Cardiovascular & Ablation Technologies, St Jude Medical Inc.
5Radiology, Stanford University, Stanford, CA
7Siemens Healthcare R&D, Boston, MA
8Johns Hopkins University, Baltimore, MD
9MRI Interventions, Irvine, CA
10Sunnybrook Health Sciences, Toronto, ON

## Slide 2
Declaration of Relevant Financial Interests or Relationships
Speaker Name: Ehud J Schmidt
Presented MR-tracking and MRI-compatible VDT-tracking systems are not FDA Regulatory Cleared.
Patents submitted for (1) MR Tracking on Metal surfaces and (2) MRI-compatible Electro-Anatomic Mapping Systems
Jeff Schweitzer and Israel Byrd are employees of St. Jude Medical Inc.
Ravi Seethamraju is a Siemens Healthcare employee
Kamal Vij is an employee of MRI Interventions Inc.
Ehud Schmidt receives grant support from St. Jude Medical, E-TROLZ
Research Support: NIH U41-RR019703,NIH R43-HL110427-01, AHA 10SDG261039, NIH U54-HL119145, NIH-R21CA167800
Notes: 2

## Slide 3
New developments in MR-Tracking
Phase Dithering
Tracking on Metallic Surfaces
Micro-transmit
Directionality detection using phase
Voltage Device Tracking
Technique
First implementations in MRI
Imaging while tracking
3
Topics

## Slide 4
RF
pulse
Acquisition
Freq. Encoding
Gradient
Large Structure
Micro-coil
Freq. Encoding Grad.
Dephasing Grad.
Dephasing Gradient
Dumoulin et al. Magn Reson Med 2010;63:1398-1403;
Reduce the effect of
Coupling of surrounding coils or catheters
B1 in-homogeneities
Combine the data acquired over the cycle
Improved peak, robust MR tracking
MR-Tracking with Phase Dithering

## Slide 5
3 MRI Devices Tracked
Simultaneously
Deflectable Sheath
Deflectable RF Ablation catheter
Non-deflectable Coronary Sinus catheter
Application of Phase Dithering reduces device coupling & flow artifacts
Cardiac EP using active MR-Tracking
SR Dukkipati et al. Circulation 2008;118: 853-862
EJ Schmidt et al. Circ. Arrhythmia Electrophysiol. 2009;2:695-704

## Slide 6
MR-Tracking of Metallic Devices
Most clinical intv’l devices (guidewires, catheter, sheathes, needles)
are largely made of metal
Improved mechanical properties for narrow and long devices
(larger stiffness, higher yield point, reduced brittle or fluid behavior,
better torquability)
Current MRI-compatible devices use a minimum of metal
Weaker devices, larger diameter for equivalent performance.
Better tracking properties.
Can we make active metallic MRI-compatible devices?
Active Tracking on Metal is Challenging
Static Magnetic Field (B0) In-homogeneities close to metal surface
RF Field (B1) In-homogeneities
Possible RF heating of long devices

## Slide 7
MR-Tracking of Metallic Devices
Multilayer Flexible printed circuit coils on
1.6 mm OD Tungsten needle
Simulation of RF (B1 ) Field
2 mm
Sagittal MR image acquired by the micro-coil
W. Wang et al., Mag. Reson. Med. 2015; 73: 1803-11
W. Wang et al. , Medical Physics 2015; 42: 7114-7121

## Slide 8
MR-Tracking Brachytherapy Stylets (N=10)
1 cm
Active Tracked Metallic Stylet 33 cm length (MRI Interventions)
70 yr-old female
with cervical cancer
Patient Isolation
Unit
8-channel
MRI receiver
Tuning, matching, decoupling

## Slide 9
MicroTransmit (Tracking Transmit&Receive)
Coils
Magnet
Body RF coil
Transmit
Receive
Tracking coil
Body coil transmit / Tracking receive
Uniform RF excitation
High power (15KV)- device heating?
Large spin dephasing
Tracking coil transmit & Receive
Localized RF excitation
Low power (.2KV) - no device heating
Low spin dephasing
50% smaller location error
Requires hardware changes
BR Daniels et al. Mag. Reson. Imaging 2016;34: 289-297
Magnitude
Magnitude
Body coil
Transmit
Micro-
Transmit
0
12
24
36
48
0
12
24
36
48
Coils
Magnet
Body RF coil
Transmit
Receive
Tracking coil
Switch

## Slide 10: Directionality from Micro-coil Phase
MR-tracking projections do not give information on device orientation.
Two micro-coils are now required to determine direction
Phase images are unique to the position and orientation of the micro-coil.
Two low-resolution projection phase images in orthogonal planes can be used to determine device position and orientation.
Possible issues with magnetic susceptibility
Phase - plane 1 (x-z)
Phase - plane 2 (x-y)
Anderson et al. IEEE Trans. Med. Img. 31(6):1173, 2012
Notes: -Phase images acquired in two orthogonal planes contains is unique to the position and orientation of a microcoil.
-Phase images can be therefore used to determine the position and orientation of a microcoil.

## Slide 11: Tracking for Multi-modality EP Intervention
Foreseeable future
Cardiovascular Intervention performed using multiple modalitiesX-ray, Intra Cardiac Echo (ICE), MRI
Existing Tracking Methods:
MRI-based tracking works only inside MRI.
Magnetic tracking works only outside MRI
Voltage device tracking (VDT) works outside MRI & inside MRI when scanner is not pulsing.
VDT used in >80% of EP sites (St. Jude NavX-VelocityTM & Biosense-Webster CarTo3TM)
ICE showing Foramin Ovalis tenting during Trans-septal puncture
X-Ray vascular navigation
Voltage-based SJM Velocity tracking system
LA
RA
MRI showing anatomy & RF ablation gaps

## Slide 12: Voltage-Based Device Tracking: Principles*
*Wittkampf FHM et al. Circulation 1999, 99:1312-1317
Right
Left
Back
Front
Head
Feet
SysRef
8 kHz voltage bursts between 3 pairs of surface patches.
Ionic conduction through body
Localization based on voltage division across body impedance
Multiple catheters tracked @90 fps
Surface Patches Position
Slotted Surface Patches
Transmission between patches
Body as voltage divider
Receive at catheter electrodes
Right
Left
10 mV
9 8 7 6 5 4 3 2 1

## Slide 13: VDT Electro-Anatomic Mapping: MRI Setup
8 KHz Signal Bursts
Transmit to 6 Patches
SysRef
Voltage Tracking
Ground
. Back
. Left
.Right
. Head
.
. Front
. Foot
Received from catheter electrodes 1, 2,3,….N (N<64)
20 MHz RF
Low-Pass Filters
Patches
SysRef
Voltage-tracking Ground
Penetration Panel
Scanner Ground
Catheter
Channel 1
Switch
Delay
Channel 2
Switch
Delay
Channel 3
Switch
Delay
Channel N
Switch
Delay
.Channel 4
.Channel 5
.
Intra-cardiac ECG
Position (8KHz)
Intra-cardiac ECG
Position (8KHz)
Intra-cardiac ECG
Position (8KHz)
Intra-cardiac ECG
Position (8KHz)
TTL from
Grad & RF
TTL from
Grad & RF
TTL from
Grad & RF
TTL from
Grad & RF
Penetration Panel
Scanner Ground
MRI Room
Control Room
Control Room

## Slide 14: VDT Tracking in MRI during 3T imaging
Schmidt et al. Mag. Reson. Med 2014;71: 1197–1209
Irrigated Deflectable EP catheter
Display on St. Jude NavX system
Three catheters tracked at 20 fps
Tracking & mapping simultaneously

## Slide 15: EAM During & Outside MRI: Comparison
15
Catheter Tracking over 60-80% of sequence.
Locations
(i) Scan vs. No scan Mean: 0.33 0.37mm SD: 0.23 0.45mm
(i) Inside vs Outside Mean deviation <0.6mm
Current Limitations: When TR<32msec (1) TR extension (2) imaging-pauses
EAM During MRI
EAM outside MRI
EAM registered with MRA*
Location deviation
* West JJ, Cardiovasc Electrophysiol 2010; 21, 81-87

## Slide 16: VDT Catheters during MRI Imaging un-gated axial SSFP
16
Coronary
Sinus
RV
Coronary
Sinus
MR image quality reduction< 5%

## Slide 17: EAM Mapping in the Left Ventricle (JHU)

## Slide 18: 2 EPs Playing inside MRI using VDT
Schmidt et al. Mag. Reson. Med 2014;71: 1197–1209
Irrigated Deflectable EP catheter
Display on St. Jude NavX system

## Slide 19
Conclusions
Precision & Robustness MR-Tracking improvements
Ability to use metallic tracked intv’l devices (heating-resolution limited)
Tracking with Voltage Device Tracking inside and outside MRI
Ability to track with VDT during limited MR imaging
Potential Benefits:
Multi-modality intervention (X-ray, ICE, MRI) using VDT-mounted devices without device co-registration for shorter & cheaper procedures
Dual MR-Tracking/Voltage Device Tracking catheters for greater geometric accuracy VDT-based electro-anatomic mapping & ablation
Easier MRI entry into clinical Electrophysiology (cheaper catheters, VDT hardware & software available at most EP sites)
19
Conclusions