## Slide 1: Progress MRI Guided Adult Rt Heart Catheterization
Alexander (Sandy) Dick, MD
Rebecca Thornhill, PhD
Sophie Chenier, MRT(R,MR)

## Slide 2: Summary
iCMR @ NHLBI
  NHLBI Hands-On Workshop September 28, 2016
  Ratnayaka et al, Eur Hrt J 2012; Rogers et al, CHEST, 2014
Progress to date OHI
  IRB
  Equipment
  Safety

## Slide 3: Challenges
Justification
Cost/Access
  Limited vendors in small market
Stand alone busy clinical CMR
Safety

## Slide 4: Justification MRI Guided Catheterization
Simultaneous structural, functional and hemodynamic assessment
  RV afterload, contractility and right ventricle-pulmonary artery coupling
Diagnosis and prognosis
Future MRI guided procedures

## Slide 5
Example of Post contrast T1, partition coefficient and ECV maps of a 68 year-old female with severe chronic thromboembolic pulmonary hypertension, with an mPAP=72 mmHg on right heart catheterization.
Example of Post contrast T1, partition coefficient and ECV maps of a 68 year-old female with severe chronic thromboembolic pulmonary hypertension, with an mPAP=72 mmHg on right heart catheterization.
Peña E... Mielniczuk L. SCMR 2016.

## Slide 6: Interventional MRI Suite Equipment List
Item | Mfr | Model | Qty | Cost | Total
Projection System |  |  |  |  | 
Projector | InFocus | IN3120HD | 2 | $1500 | $3000
Screen | Da-Lite |  | 1 | $700 | $700
Video Matrix Switch | NTI | SM-16X16-DVI | 1 | $3800 | $3800
Fiber optic extender | NTI | ST-2FODVI-LC | 4 | $600 | $2400
Fiber Cable | NTI | FIBER-2D-LCLC | 4 | $170 | $680
XFM |  |  |  |  | 
PC | Dell |  | 1 | $4000 | $4000
Graphics card | Nvidia | GTX970 | 1 | $600 | $600
A/D card | National Instruments | NIUSB-6009 | 1 | $600 | $600
PHYSIO-recording |  |  |  |  | 
Invivo and MACLAB |  |  |  | $0 | $0
Electronic Filters | In House |  | 1 | $3000 | $3000
Communication |  |  |  |  | 
Wireless headsets (4) | Innovere |  | 1 | $35000 | $35000
TOTAL |  |  |  |  | $53780

## Slide 7: Wireless Communication

## Slide 8: Optoacoustics
Hands free MRI noise cancelation communications system
IMROC™ Wireless: Complete Freedom and Flexibility
x 6

## Slide 9: R.F. Shielded Projector

## Slide 10: The NIH Lab at Children’s National Medical Center Washington, DC
Notes: Standard 1.5T magnet….
- Avoid 3T d/t SAR issues which are problematic when using intravascular devices.
- Avoid banding artifacts common with SSFP imaging techniques.
Intermodality transfer table…
- For procedures requiring hybrid imaging
- For bailout from MR to X-ray in emergent situations

## Slide 12
Control Room

## Slide 13: Data Display NHLBI

## Slide 15: Safety (What keeps me up at night)
Access in the room
  Dedicated removal of needle and wire
MRI safe devices
Arrhythmias
  Complete Heart Block, Ventricular Tachycardia,Ventricular Fibrillation
  Dedicated observation monitor; arterial access

## Slide 16: MR Safe Procedure Table

## Slide 17: Patient Evacuation
 | 1st Operator | 2nd Operator | Cath/MR Tech | Cath RN | Sedation RN
1 |  | Provide Instructions |  |  | 
2 | Advance patient to scanner home position |  | Open MR bay doors |  | Secure IV line
3 | Remove Surface Coil |  | Remove pt headset |  | Secure airway
4 | Start chest compressions |  | Transfer patient from MR suite |  | Assist Cath RN
5 |  |  | Restricts access to MR suite | Defibrillate/ Administer medications | 

## Slide 18: Summary
iCMR @ NHLBI
Progress to date OHI
  IRB
  Equipment
  Safety