## Slide 1: EP instrumentation/devices
Bruno Quesson
IHU LIRYC & Inserm U1045
University of Bordeaux - France

## Slide 2: Contact electrophysiology
Catheters
3D Navigation system
Electromagnetic 3D positioning
EP console
RF ablation device
3D electroanatomical mapping
+
=

## Slide 3: Setup for MR-EP
MR Console
(SIEMENS)
Acquisition/Reco
MRI
EP console
Electrophysiology
RFA
generator
Milli V
~1 Hz
80 V
0.5 MHz
> 100 V
64 MHz

## Slide 4: Inside the MR room
Source: usa.healthcare.siemens.com
1.5T MR scanner
(SIEMENS)
Catheter
F
F
RFA
(Stockert)
EP

## Slide 5: Safety and artifacts
Heating test
in Agar gel
(0.9% saline)
Elbes et al, Europace. 2016
In vivo EP in sheep
Catheter perpendicular to B0
Catheter parallel to B0
Catheter design
(Material of the electrode)
Catheter A
Catheter B

## Slide 6: Improving EP during MRI scanning
Zhang SH et al, Magnetic Resonance in Medicine 75:2204–2216 (2016)
1 - acquire EP signals before MR scanning
2- acquire EP signals during scanning
3- subtract => gradient induced voltages
4- retrieve gradient amplitudes and time derivatives
5- fit data to F to derive the 19 coefficients
Pertubation = F(Gx, Gy, Gz, dGx/dt, dGy/dt, dGz/dt )
12 leads ECG
19 coefficients to determine
collect gradient waveforms and use the coefficients to correct ECG
traces during any MRI scanning sequence
Learning step
During scanning

## Slide 7: MRI : 3D catheter navigation
Electrodes
MR compatible catheter
Micro-coils
Dumoulin CL, et al. Real-time position monitoring of invasive devices using magnetic resonance, MRM, 1993
Vergara GR, et al. Real-time magnetic resonance imaging-guided radiofrequency atrial ablation and visualization of lesion formation at 3 Tesla, Heart Rythm, 2011
Wech T. et al. Measurement accuracy of different active tracking sequences for interventional MRI, Journal of Magn. Res. Imag., 2013
RF
X
Y
Z
ADC
X
Y
Z
Imaging sequence
Micro-coil position detection
Slice position update
SAG
Tip tracking slice following sequence
BEAT-IRTTT (Interactive Real-Time Tip Tracking) pulse sequence
Bssfp interleaved with active tracking module
1.8 x 1.8 x 4 mm – FOV: 225 x 225 mm
Frame rate: 3fps
3 mm
In vivo in sheep

## Slide 8: MR-Thermometry
Slice #1
Slice #2
Slice #3
Slice #4
Temperature elevation (°C)
Temperature elevation (°C)
Cumulative thermal dose (°C)
Lethal threshold
Hand-drawn ROI in post-processing
Toupin el al, JCMR 2017

## Slide 9: Post-ablation lesion assessment
Voxel size 1.3x1.3x2.5 mm3
31 segments, 60 slices
Two-beats ECG triggered
Respiratory gated
TI = 700 ms
IR-3D sequence
Post-ablation visualization of all RF lesions at once in 3D
No contrast agent administration (Gadolinium)
Contrast-to-noise ratio (CNR) = 88.8 ± 38.3
Lesion
“Towards MRI-guided cardiac ablation procedures with no contrast agent”, Guttman M, Susumu T, Kolandaivelu A et al., SCMR 2015

## Slide 10: conclusion
Dedicated Hardware
  Catheters
  Filters
  Navigation system
  Defibrillators
  … (sensors)
Safety and artifact free
  No heating of devices
  Limited image artifact
  EP during MRI scanning
MRI
  Structural imaging (refine target definition), integration into EP consoles (therapy planning)
  Real-time MR-thermometry (control the procedure)
  Post-ablation lesion assessmentg (check treatment)

## Slide 11
11
IHU Liryc
NIH
Valery Ozenne
Solenn Toupin
Pierre Bour
Fabrice Marquet
Thibaud Troadec
Manuel Boissenin
Julie Magat
Jerôme Naulin
Mathilde Merle
Florent Collot
Pierre Jaïs
Arnaud Chaumeil
Fanny Vaillant
Virginie Loyer
Delphine Vieillot
Marion Constantin
… all IHU LIRYC!
Michael S. Hansen
Souheil Inati
Financial support:
ANR 11-Tecs-003 TACIT
ANR 11-EQPX0030 Equipex MUSIC
ANR 10-IAHU-04 IHU LIRYC
France Life Imaging
Baudouin Denis de Senneville
IMB
Matthieu Lepetit-Coiffé
Siemens
Jenny Benois-Pineau
Pascal Desbarats
Aurélie Emilien
LABRI