## Slide 1: Interventional CMR using off-the-shelf devices
Dr. Toby Rogers, MA, BM BCh, PhD
Dr. Adrienne Campbell-Washburn, PhD
Cardiovascular Branch, Division of Intramural Research National Heart, Lung, and Blood Institute National Institutes of Health Bethesda, MD, United States

## Slide 2: Acknowledgements
Robert Lederman, MD
Kanishka Ratnayaka, MD
Daniel Herzka, MD
William Schenke, BA
Jonathan Mazal, MA
Jaffar Khan, MD
Elena Grant, MD
Adrienne Campbell-Washburn, PhD
Raj Ramasawmy, PhD
Peter Kellman, PhD
Laurie Grant, RN
Annette Stine, RN
Merdim Sonmez, PhD
Ozgur Kocaturk, PhD
Michael Hansen, PhD
Marcus Chen, MD
Burcu Basar, BS
Anthony Faranesh, PhD

## Slide 4
X-ray guided interventions
Image courtesy of Bill Schenke and Delaney McGuirt

## Slide 5
X-ray guided interventions
Image courtesy of Bill Schenke and Delaney McGuirt

## Slide 6
iCMR guided interventions
Image courtesy of Bill Schenke and Delaney McGuirt

## Slide 7
Device safety
Heats during imaging

## Slide 8
iCMR guided interventions
Image courtesy of Bill Schenke and Delaney McGuirt

## Slide 9: Non-metallic catheters
Cook, Flow Directed Balloon Catheter, Model #: FDB5.3-35-80
Medtronic, Pulmonary Wedge Pressure Catheter, Model #: 150075
Vascor, Balloon Wedge Pressure Catheter, Model #: 172-110P
Edwards, True Size Monitoring Catheter, Model #: 111F7
Edwards, True Size Hi-Shore Monitoring "T" Tip Catheter, Model #: T111F7
Arrow, Balloon Wedge Pressure Catheter, Model #: AI-07127
Edwards, True Size Monitoring “S” Tip Catheter, Model #: S111F7

## Slide 12: MR fluoroscopy guided cardiac catheterization

## Slide 13
Standard MR Fluoroscopy
ΔT = 18°C
ΔT = 11°C
ΔT = 2.3°C
ΔT = 0.22°C
Spiral MR Fluoroscopy
ΔT = 0.15°C
ΔT (°C)
2 min MR fluoroscopy
Heating is related to imaging parameters
time
0.035” 145cm Nitinol guidewire (Nitrex, Medtronic)
70x reduction
Image courtesy of Adrienne Campbell-Washburn, PhD
Notes: Guidewire type

## Slide 14
Not all metallic guidewires are the same

## Slide 15: Clinical trial
Methods
7 patients
46±20 years
4 (57%) women
Clinically-indicated MR fluoroscopy right heart catheterization at the NIH Clinical Center
0.035” 150cm angled-tip Terumo Glidewire and Edwards T-tip balloon end-hole catheter

## Slide 16: Clinical trial
Endpoints
Procedural success
Guidewire and catheter conspicuity
Adverse events
Pre- and post-procedure D-dimer and plasma free Hgb

## Slide 17: Clinical trial
Results
100% procedural success
Guidewire imparts incremental stiffness and conspicuity to the catheter shaft
Guidewire tip often not visible
No adverse events
No change in blood assays

## Slide 18
Off-the-shelf 0.035” guidewire (Terumo)
Spiral gradient echo
Low flip angle (10)
Long TR spiral readout (10msec)
MR fluoroscopy IVC  SVC

## Slide 19
iCMR guided interventions
Image courtesy of Bill Schenke and Delaney McGuirt

## Slide 20
iCMR guided interventions
Image courtesy of Bill Schenke and Delaney McGuirt

## Slide 21: Summary
Most X-ray catheters and guidewires heat during standard MR fluoroscopy
Engineering dedicated devices for MRI is hard
Engineering devices for MRI with the required mechanical properties to perform interventions is even harder

## Slide 22: Summary
First-in-human report of cardiac catheterization with a commercially available metallic guidewire
Low-SAR MR fluoroscopy enables select metallic guidewires to be used safely for MR interventions
Modifying imaging parameters rather than re-engineering devices will accelerate clinical translation of MR interventions

## Slide 23: CAUTION!