## Slide 1: MR guided electrophysiology imaging and catheters
Solenn TOUPIN, PhD
Electrophysiology and cardiac modeling institute – Bordeaux, France

## Slide 2: Disclosure
Employee of Siemens Healthineers France

## Slide 3: MR guided electrophysiology imaging and catheters
Catheter guidance
Ablation monitoring
Post-ablation lesion
imaging

## Slide 4: Catheter guidance : passive tracking
Courtesy of Adrienne Campbell
b-SSFP with positive contrast
Through-place dephasing for guidewire visualization
Positive contrast spiral imaging for visualization of commercial nitinol guidewires with reduced heating, A. E. Campbell-Washburn, T. Rogers et al, JCMR 2015
Improved passive catheter tracking with positive contrast for CMR-guided cardiac catheterization using partial saturation (pSAT), M Forte, K Pushparajah et al, JCMR 2017
Balloon catheter filled with Gd
b-SSFP with partial SAT
Adapted from supplementary material video
Real-time MRI-guided right heart catheterization in adults using passive catheters, K Ratnayaka A Faranesh, EHJ 2013
Balloon catheter filled with air
b-SSFP

## Slide 5: Catheter guidance : active tracking
Micro-coils
B0
Frequency peak
Magnetic field gradient

## Slide 6: Interactive interface
Compatible with active and passive tracking
3D volume and segmentation of heart chambers
Real-time interactive sequence (b-SSFP)
Electro-anatomical mapping if coupled to an EP software
Markers for mapping and ablation spots
Gaspar et al., EHJ 2014
Epicardial electroanatomical mapping, radiofrequency ablation, and lesion imaging in the porcine left ventricle under real-time magnetic resonance imaging guidance—an in vivo feasibility study, R. Mukherjee, S. Roujol et al, JCMR 2017

## Slide 7: Ablation monitoring: real-time MR thermometry
Challenges for iCMR :
High temporal resolution ~1s
High spatial resolution
Motion correction between phase maps
Inline calculation of MR thermometry
Temperature precision < 2°C

## Slide 8: Ablation monitoring : real-time MR thermometry
Slice #1
Slice #2
Slice #3
Slice #4
Temperature (°C)
Cumulative thermal dose
Lethal threshold
Temperature (°C)
Feasibility of real-time MR thermal dose mapping for predicting radiofrequency ablation outcome in the myocardium in vivo, S. Toupin, P. Bour et al JCMR 2017

## Slide 9: Ablation monitoring : real-time MR thermometry
Temperature evolution (°C) in 5x5 voxels
Time (s)
Feasibility of real-time MR thermal dose mapping for predicting radiofrequency ablation outcome in the myocardium in vivo, S. Toupin, P. Bour et al, JCMR 2017

## Slide 10: Ablation monitoring : MR dosimetry
Epicardial electroanatomical mapping, radiofrequency ablation, and lesion imaging in the porcine left ventricle under real-time magnetic resonance imaging guidance—an in vivo feasibility study, R. Mukherjee, S. Roujol et al, JCMR 2017
Feasibility of real-time MR thermal dose mapping for predicting radiofrequency ablation outcome in the myocardium in vivo, S. Toupin, P. Bour et al, JCMR 2017
Thermal dose
Thermal dose
T1-w imaging
Gross macroscopy

## Slide 11: Ablation monitoring : chemoablation
Real-time visualization of ablation lesions
Irreversible necrosis
Transmural lesions

## Slide 12: Post-ablation visualization of lesion
Late gadolinium enhancement
15 min after injection
85 min after injection
Gross pathology
Adapted from Figure 1
Characterization of radiofrequency ablation lesions with gadolinium-enhanced cardiovascular magnetic resonance imaging, T Dickfeld, R Kato, J Am Coll Cardio, 2006
Vijayakumar et al.: Assessment of cardiac RF ablation lesions by DCE-MRI. Journal of Cardiovascular Magnetic Resonance 2014

## Slide 13: Post-ablation visualization of lesion
Non contrast T1 weighted volume: TI 700-800 ms
Post-ablation visualization of all RF lesions at once in 3D
Absence of contrast agent
Voxel size 1.1x1.1x2.5 mm3
2 RR ECG triggered
Respiratory gated
T1w (MPR SAX)
Gross Photography
T1w
Gross Photography
M: Normal myocardium
CBN: Contraction band
necrosis / hemorrhage
CN: Coagulation necrosis
Thrm: Thrombus
Courtesy of Michael Guttman
Non-Contrast-Enhanced T1-Weighted MRI of Myocardial Radiofrequency Ablation Lesions
M. Guttman, S Tao et al, MRM 2017

## Slide 14: Acknowlegments
Bruno Quesson
Valery Ozenne
Pierre Bour
Slides
Christopher Piorkowski
Adrienne Campbell
Michael Guttman
Rainer Schneider
Thomas Pohl