## Slide 1
ICMR Defibrillators
Ehud J Schmidt1
JHU/BWH/Stanford prototype: Ehud J Schmidt1,4, Ronald D Watkins2, Menekhem M Zviman1,3, Michael A Guttman1, Wei Wang4, Hassan Elahi1, Eric S Meyer1, Ron Berger1, Henry R Halperin1
PinMed/Iowa/NHLBI prototype: V. Shusterman5,6, Denice Hodgson-Zingman6, Daniel Thedens6 , Xiaodong Zhu6 , Stacy Hoffman5 , Jessica Sieren6 , Gina M. Morgan6 , Anthony Faranesh7 , Rajiv Ramasawmy 7 , Adrienne Campbell-Washburn3, Barry London6
1Cardiology, Johns Hopkins University, Baltimore MD. 2 Radiology,
Stanford University, Stanford CA, 3University of Pennsylvania, Philadelphia PA
4Radiology, Brigham and Womens’ Hospital, Boston MA
5PinMed, Inc., Pittsburgh PA, 6The University of Iowa, Iowa City IA,
7NHLBI, Bethesda, MD

## Slide 2
Declaration of Relevant Financial Interests or Relationships
Speaker Name: Ehud J Schmidt
Presented Defibrillation and 12-lead ECG systems are not FDA Regulatory Cleared
Patents have been submitted for the MRI-compatible External Defibrillator, and the Pain-less defibrillation by Johns Hopkins University
JHU prototype Grant Support: NHLBI R01-HL094610
Pinmed prototype Grant Support: NHLBI contract HHSN268201500009C
US government employees: Anthony Faranesh, Rajiv Ramasawmy, Adrienne Campbell-Washburn
Pinmed employees: Vladimir Shusterman, Stacy Hoffman
Notes: 2

## Slide 3
Why is Defibrillation in MRI required?
Requirements for MRI-compatible defibrillation systems
System Design
System Testing
Human Volunteer safety and RFI testing
Animal defibrillation results
Addressing the Muscle contraction issue
Reducing Motion and Pain
Topics

## Slide 4
MRI Imaging/intervention on severely-ill patients is not performed
12-lead ECG monitoring of patients not commercially available.
without 12-lead ECG, detection of acute ischemia may be very late!
Defibrillation now requires;
  Patient removal from magnet bore
  Coils and other peripherals disconnected
  Patient removal from MRI table and MRI suite
valuable time is lost, with each minute reducing survival by 5%
Patient populations currently excluded from MRI imaging
Acute tachycardia or higher-risk post-infarct patients
Trauma (stroke, spine & orthopedic trauma)
MRI exercise stress tests on higher-risk patients
Patient populations currently excluded from non-cardiovascular
Or cardiovascular MRI-guided intervention
Unstable (Blood Pressure, Heart Rate) patients
Patients with implanted devices*
Patients with ischemic histories
Need for Defibrillation in MRI Suite

## Slide 5: Ventricular Tachycardia during EP Atrial Fibrillation ablation 2 hrs into procedure =>MRI Table undocked for Defibrillation.
Swine Example of Late Detection of VT in a 3T MRI study

## Slide 6
Defibrillation & CPR performed in the MRI Room
Defibrillation is frequently performed during
Human EP (VT) interventions!
Swine Example of Late Detection of VT in a 3T MRI study

## Slide 7
7
Contrast MRA
T2
Balloon in LAD
Inflated to 20 Atm.
20 min until death
Continuous ECG
Continuous MRI
Tse ZTE et. al MRM 2013, Circulation 2012
Detection of Acute Ischemia in a 1.5 T MRI with a 12-lead ECG

## Slide 8
On demand (2 KV) defibrillation in the MRI Magnet
Defib. Generator (or control) in the MRI Room
Pads continuously on patient & connected to generator
Generator always ON, in ECG monitoring mode
Generator senses ECG pads (Zoll: 67 KHz) - ready for Defib delivery
Unhindered MR Imaging performance
Minimal SNR loss due to connection of defibrillator and pads
Minor Image artifacts due to defibrillation pads
RF heating <1.5 Co @ 4 Watts/kg SAR
MRI system and peripherals withstand defibrillation pulses
Patient monitoring system
12-lead ECG monitoring system
MR Imaging equipment (MRI electronics, MRI coils)
Requirements for an MRI-compatible Defibrillator

## Slide 9
MRI-Compatible Defibrillator Block Diagram
Zoll Medical
M-Series
Defibrillator
Generator
10-Pole Low-Pass
Filter
Twisted Pair
High-voltage
Cable w/ RF Baluns
every 30 cm
Zoll Stat Padz
Multifunction
Defibrillation
Pads
Differential Mode
Unit
Common Mode
Unit
Low Pass
Filter
Unit
Defibrillator
Generator
RF Baluns
High Voltage
Twisted Pair
Cable
Defib. Pads
(Disposable)
MRI-Compatible Defibrillator Components (JHU system)

## Slide 10: Human MRI SNR and Safety testing
Defibrillation Pad
Zoll M
Defibrillator
Generator
Baluns
LP Filter
System Layout In MRI Suite
System OFF
System ON
Battery power
System ON
A/C Line power
SSFP
Slice 1
SSFP
Slice 2
pad artifact
Defib Pad artifacts extend only 6 mm from surface (SSFP)
SNR Reduction of 1.6 dB=17% (A/C power) , 0.8 dB=8% (battery power)
~90dB reduction in 63.8 MHz noise
No RF heating during imaging

## Slide 11
Hemodynamic Display during cardiac arrest
Hemodynamic Display before Fibrillation
Hemodynamic Display after
Defibrillation
Swine in 1.5T MRI during cardiac arrest
Zoll Generator and filter in MRI room
Defibrillation
Cable
Swine Fibrillation and Defibrillation Testing (II)

## Slide 12
Medtronic LifePak20defibrillator
custom cables (Pinmed)
12 lead Mri compatible ECG (Pinmed)
1.5 T MRI
Defibrillator pads in the left-posterior axillary (1) and right subclavian (2).
Temperature under defibrillator pads recorded using fiber-optic temperature probes (arrows).
18 experiments.
Defibrillation inside bore during continuous 2D SSFP imaging.
Defibrillation Experiments (Pinmed)
Shusterman V et. al, Circulation 2017;136:A13735

## Slide 13
Defibrillation successful in 100% of experiments at 1.5 T
10% reduction in SNR with Defibrillation Generator ON.
ECG in lead V1 before defibrillation and after defibrillation during MRI scanning (SSFP) in a swine;
Fibrillation
Post-Defibrillation
Time [sec]
ECG [mV]
Time [sec]
Time [sec]
ECG [mV]
ECG [mV]
Swine Defibrillation Success (Pinmed)
Shusterman V et. al, Circulation 2017;136:A13735

## Slide 14
Temp. changes before, during and
after defib
Temperature 10 sec after defibrillation
Max
Temperature changes during defibrillation (Pinmed)

## Slide 15: MRI Defibrillation in action (small swine subject)
Note:
1. jump has nothing to do with MRI
Magnetic field !!
Magnetic (Lorentz) accel. (3T) ~0.33ms-2
Levitation accel. (1g)=9.8ms-2
The Jump you see is >7g=70ms-2
2. Completely due to abdominal
muscle contraction.

## Slide 16: Defibrillation without Muscle Contraction
Hunter DW et. al, Heart Rhythm 2016;13(5): 1142

## Slide 17: Defibrillation with minimal Muscle contraction
Modified Defibrillation Waveform
Preparation: 1 kHz waveform, ramped over 1s from 0->25V
Biphasic Voltage: 480 V for 4 ms per phase (9J)
71% mean reduction in acceleration (to ~4.7g)
72% mean reduction in force (pain)
Same success in swine defibrillation as with Biphasic
Hunter DW et. al, Heart Rhythm 2016:13(5) 1142
levitation

## Slide 18: Defibrillation in the MRI
MRI-compatible external defibrillators
Constructed by adding Radio-Frequency filters and changing cables.
Patient imaged with pads on chest and defibrillator ON in monitoring mode (charged and measuring ECGs) to shorten response time
Insignificant imaging SNR penalty (~7%)
No patient (pad) heating risk due to MRI RF pulses
Proper function of defibrillator inside the bore.
Motion during defibrillation reduced by modifying defib waveform
Thank You

## Slide 19
Publications
Schmidt EJ, Watkins RD, et al. A magnetic resonance imaging–conditional external cardiac defibrillator for resuscitation within the magnetic resonance imaging scanner bore. Circ Cardiovasc Imaging. 2016;9:e005091. doi:
Shusterman V, Hodgson-Zingman D, et al. External Defibrillation During MRI: Feasibility and Safety Circulation. 2017;136:A13735
Hunter DW, Tandri H, et al. Tetanizing pre-pulse: a novel strategy to mitigate ICD shock-related pain Heart Rhythm. 2016 May ; 13(5): 1142–1148.